DESIGN CONTROL PROCEDURE

This procedure defines the design and development control process for medical devices, accessories, packaging, labeling, production processes, software, and significant product changes. It is intended for a large organization operating multiple design centers, manufacturing sites, suppliers, and commercial regions, and is structured for implementation within an ISO 13485 quality management system.

# DOCUMENT CONTROL INFORMATION

| Field | Required entry |
| --- | --- |
| Document ID | STK-ISO13485-P4-SOP-001 |
| Process area | ISO 13485 quality management system - design and development control |
| Document owner | Director, Design Assurance |
| Approved by | Chief Quality Officer |
| Effective date | [YYYY-MM-DD] |
| Version | 1.0 |
| Confidentiality | Internal controlled document |
| Review cycle | Annual or upon significant regulatory, process, product, or organizational change |

[Note] for implementation consultants: Confirm that this document is issued through the organization document control procedure and that the effective version is available at points of use across design, manufacturing, regulatory, purchasing, service, and commercial sites.

## REVISION HISTORY

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Version | Date | Description of change | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Version | Date | Description of change | Prepared by | Reviewed by | Approved by |
| --- | --- | --- | --- | --- | --- |
| 1.0 | 2026-01-15 | Initial release aligned with ISO 13485:2016 clauses 4, 5, 6, 7 and 8. | Elena Fischer | Martin Keller | Sophie Laurent |
| 1.1 | 2026-03-02 | Added design transfer governance and post-market feedback escalation. | Thomas Berger | Charlotte Meyer | Sophie Laurent |
| 1.2 | 2026-05-20 | Updated risk management references to ISO 14971:2019. | Amelia Ross | Lukas Schneider | Isabelle Moreau |
| 1.3 | 2026-07-18 | Expanded supplier interface controls for outsourced design verification. | Nicolas Weber | Eva Hartmann | Sophie Laurent |
| 1.4 | 2026-09-10 | Added EU MDR and UDI linkage in quality planning records. | Marta Klein | Oliver Novak | Isabelle Moreau |
| 1.5 | 2026-11-29 | Included global distribution quality requirements and recall communication matrix. | Julian Meyer | Anne Dupont | Sophie Laurent |

# 1. PURPOSE

The purpose of this procedure is to ensure that medical device design and development activities are planned, controlled, reviewed, verified, validated, transferred, changed, and documented in a manner that demonstrates conformity with design inputs, user needs, intended use, applicable regulatory requirements, and risk management requirements. The procedure establishes a consistent stage-gate framework that integrates design controls with risk management, usability engineering, software life-cycle controls, supplier management, production readiness, design transfer, post-market feedback, and design change control.

[Note] for implementation consultants: Design control procedures should be linked directly to the device master record, design history file, risk management file, usability file, software documentation, clinical evaluation, labeling, production validation, and regulatory submission strategy.

# 2. SCOPE

This procedure applies to new product development, platform extensions, product redesigns, significant design changes, manufacturing process designs that affect device specifications, design transfers, software changes, packaging and labeling design, outsourced design activities, and design changes initiated by complaints, CAPA, field actions, regulatory changes, supplier changes, technology changes, or post-market surveillance findings.

Class I, II, III and in vitro diagnostic devices where the organization has design responsibility or regulatory accountability.

Hardware, firmware, software as a medical device, embedded software, accessories, sterile barrier systems, packaging, labeling, instructions for use, and user interfaces.

Internal design centers, acquired product lines, co-development partners, contract design organizations, test laboratories, tooling suppliers, and manufacturing transfer recipients.

Legacy products when design changes, remediation, technical file updates, or post-market signals require controlled design activity.

# 3. NORMATIVE AND PROCEDURAL REFERENCES

ISO 13485:2016, clause 7.3, design and development.

ISO 14971:2019, application of risk management to medical devices.

IEC 62304, medical device software software life-cycle processes, when software is part of the device or is the device.

IEC 62366-1, application of usability engineering to medical devices, when user interface and use-related risk are applicable.

Applicable product standards, regulatory requirements, essential principles, general safety and performance requirements, country-specific requirements, and internal quality system procedures.

# 4. TERMS AND DEFINITIONS

## DESIGN CONTROL GLOSSARY

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Term | Definition | Typical record | Process owner | Related file | Remarks |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Term | Definition | Typical record | Process owner | Related file | Remarks |
| --- | --- | --- | --- | --- | --- |
| Design input | Physical, functional, performance, safety, usability, regulatory, and interface requirements that define the device. | Design input specification | Systems engineering | DHF | Must be measurable or verifiable. |
| Design output | Design results such as drawings, specifications, code, procedures, labeling, and manufacturing instructions. | Output specification | R&D engineering | DMR | Must enable verification against input. |
| Design review | Documented, systematic evaluation of design results at planned stages. | Design review minutes | Design assurance | DHF | Requires independent reviewer where appropriate. |
| Design verification | Confirmation that design outputs meet design input requirements. | Verification protocol/report | Verification lead | DHF | Can include test, analysis, inspection, or demonstration. |
| Design validation | Confirmation that device conforms to user needs and intended use. | Validation protocol/report | Clinical/user validation lead | DHF | Performed on initial production units or equivalents. |
| Design transfer | Controlled conversion of design outputs into production specifications and routine manufacturing controls. | Transfer plan/report | Manufacturing engineering | DMR/DHR | Must verify production capability and acceptance criteria. |

# 5. RESPONSIBILITIES AND AUTHORITIES

The cross-functional project team shall include representatives from design engineering, design assurance, regulatory affairs, risk management, clinical or medical affairs, manufacturing engineering, supplier quality, service, post-market surveillance, labeling, packaging, software quality, commercial quality, and operations as applicable. Responsibilities shall be assigned in the project plan and maintained through the design and development file.

## DESIGN CONTROL RESPONSIBILITY MATRIX

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Role | Responsibilities | Authority | Required independence | Key records | Escalation criteria |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Role | Responsibilities | Authority | Required independence | Key records | Escalation criteria |
| --- | --- | --- | --- | --- | --- |
| Project Sponsor | Approve business and resource commitment. | Approve project continuation or cancellation. | Independent from technical execution. | Portfolio approval record. | Resource constraint affecting compliance. |
| Design Owner | Own technical solution, design outputs, and design issue resolution. | Approve technical design outputs. | May be part of project team. | Specifications and drawings. | Unresolved technical risk. |
| Design Assurance Lead | Ensure procedure compliance, DHF completeness, and phase gate readiness. | Block phase gate if objective evidence is insufficient. | Independent review authority. | DHF checklist and review minutes. | Missing verification, validation, or risk evidence. |
| Risk Management Lead | Maintain risk management plan, hazard analysis, risk controls, residual risk evaluation. | Approve risk management file readiness. | Independent where feasible. | Risk analysis and risk report. | Unacceptable residual risk. |
| Regulatory Lead | Define regulatory pathway, standards strategy, and submission requirements. | Approve regulatory deliverables and claims. | Independent from design decision pressure. | Regulatory plan and checklist. | New or changed regulatory requirement. |
| Manufacturing Transfer Lead | Prepare production process, validation, acceptance criteria, and DMR. | Approve manufacturing readiness. | Independent production readiness review. | Transfer report and validation records. | Production process not capable. |

# 6. DESIGN AND DEVELOPMENT PLANNING

A design and development plan shall be established and maintained for each project before detailed design work begins. The plan shall define project scope, device description, intended use, regulatory markets, classification, applicable standards, responsibilities, interfaces, design stages, deliverables, review points, verification and validation strategy, risk management activities, design transfer activities, supplier activities, software activities, usability activities, clinical or performance evaluation, resources, and criteria for phase gate progression.

The plan shall be updated as the project evolves. Changes to the plan shall be reviewed and approved by project leadership, design assurance, and regulatory affairs when the change affects compliance, risk, claims, verification, validation, submission strategy, or launch readiness.

## DESIGN AND DEVELOPMENT PLAN REGISTER

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Project / device | Design stage | Required deliverables | Responsible owner | Gate criteria | Approval record |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Project / device | Design stage | Required deliverables | Responsible owner | Gate criteria | Approval record |
| --- | --- | --- | --- | --- | --- |
| CardioWave ECG monitor | Concept | User needs, intended use, regulatory pathway. | Design Owner | Business case approved and risks acceptable. | Gate 0 minutes |
| CardioWave ECG monitor | Planning | Design plan, risk plan, standards matrix. | Project Manager | Plan approved and resources assigned. | Gate 1 minutes |
| CardioWave ECG monitor | Input | Design input specification and traceability. | Systems Engineer | Inputs complete, clear, measurable. | Gate 2 minutes |
| CardioWave ECG monitor | Output | Drawings, BOM, software architecture, labeling drafts. | R&D Engineering | Outputs controlled and traceable. | Gate 3 minutes |
| CardioWave ECG monitor | V&V | Verification and validation protocols/reports. | V&V Lead | Acceptance criteria met or justified. | Gate 4 minutes |
| CardioWave ECG monitor | Transfer | DMR, process validation, training, release criteria. | Manufacturing Engineering | Production readiness demonstrated. | Gate 5 minutes |

[Note] for implementation consultants: The design plan is a living controlled document. Auditors expect the plan to match actual project execution, not merely describe an ideal process after the project is complete.

# 7. DESIGN INPUTS

Design inputs shall be documented, reviewed, approved, and controlled. Inputs shall address intended use, user needs, functional requirements, performance requirements, physical characteristics, interface requirements, safety requirements, risk control requirements, usability requirements, cybersecurity requirements, software requirements, biocompatibility requirements, sterilization requirements, packaging requirements, labeling requirements, production and service requirements, regulatory requirements, applicable standards, shelf-life requirements, environmental conditions, and distribution requirements.

Each design input shall be clear, complete, unambiguous, feasible, traceable, and verifiable. Conflicting, incomplete, or unverifiable requirements shall be resolved before approval or formally accepted with documented rationale and mitigation actions. Risk control measures identified through risk management shall be incorporated into design inputs or linked to appropriate design outputs and verification activities.

## DESIGN INPUT REQUIREMENTS REGISTER

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Input ID | Requirement statement | Source | Acceptance criterion | Risk link | Verification method |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Input ID | Requirement statement | Source | Acceptance criterion | Risk link | Verification method |
| --- | --- | --- | --- | --- | --- |
| DI-001 | The ECG monitor shall measure heart rate between 30 and 240 bpm. | Clinical user need | Accuracy within +/- 3 bpm. | HZ-001 | Bench performance test |
| DI-002 | The enclosure shall withstand a 1.0 m drop onto a hard surface. | Use environment analysis | No hazardous damage after drop. | HZ-014 | Mechanical drop test |
| DI-003 | The device shall alert the user when battery capacity is below 15%. | Risk control requirement | Visual and audible alarm within 10 seconds. | RC-022 | System verification test |
| DI-004 | The user interface shall support setup of a new patient profile within 90 seconds. | Usability engineering | >= 90% of representative users complete task. | UR-005 | Summative usability validation |
| DI-005 | The software shall log critical alarm events with date and time stamp. | Regulatory and service need | Log contains event, timestamp, user action. | CY-008 | Software verification test |
| DI-006 | The sterile accessory pouch shall maintain sterile barrier integrity through labeled shelf life. | Packaging requirement | No seal failure after aging and distribution simulation. | PK-003 | Packaging validation |

[Note] for implementation consultants: Design inputs are often the root cause of later verification failures. Ensure requirements are testable and that each risk control has traceable implementation and verification evidence.

# 8. DESIGN OUTPUTS

Design outputs shall be documented in a form that enables verification against design inputs and provides information needed for purchasing, production, service, installation, inspection, packaging, labeling, distribution, and acceptance activities. Outputs shall include or reference drawings, specifications, material requirements, bill of materials, software code and build records, architecture documents, labeling, inspection procedures, test methods, manufacturing instructions, special process parameters, acceptance criteria, service procedures, packaging specifications, and release specifications.

Design outputs that are essential for the proper functioning of the device shall be identified. Outputs shall be reviewed and approved before release. Output changes shall be controlled through design change control and configuration management.

## DESIGN OUTPUT REGISTER

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Output ID | Output description | Input trace | Configuration status | Release location | Approver |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Output ID | Output description | Input trace | Configuration status | Release location | Approver |
| --- | --- | --- | --- | --- | --- |
| DO-001 | System architecture specification for ECG monitor. | DI-001, DI-003, DI-005 | Released Rev B | PLM-ARCH-ECG-002 | Lukas Schneider |
| DO-002 | Main PCB schematic and layout files. | DI-001, DI-002 | Released Rev C | PLM-HW-ECG-011 | Markus Vogel |
| DO-003 | Embedded software build package v3.2.1. | DI-003, DI-005 | Released build | SW repository tag ECG_3.2.1 | Anna Richter |
| DO-004 | Enclosure drawing and material specification. | DI-002 | Released Rev D | PLM-MECH-ECG-044 | Peter Novak |
| DO-005 | User manual and quick start guide. | DI-004 | Released Rev A | Labeling control system | Claire Dubois |
| DO-006 | Packaging seal specification and pouch artwork. | DI-006 | Released Rev B | Packaging DMR folder | Eva Hartmann |

# 9. DESIGN REVIEWS

Design reviews shall be planned and conducted at suitable stages to evaluate the ability of design results to meet requirements, identify problems, assess project risk, verify readiness for the next stage, and ensure cross-functional alignment. Reviews shall include representatives of functions concerned with the design stage being reviewed and at least one individual who does not have direct responsibility for the design stage under review where appropriate.

Design review records shall document the review date, participants, stage reviewed, inputs reviewed, decisions, issues, action items, owners, due dates, unresolved risks, and approval status. Open action items shall be tracked to closure and evaluated for impact on phase gate approval.

## DESIGN REVIEW RECORD TEMPLATE

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Review stage | Inputs reviewed | Key decision | Open actions | Required attendees | Approval status |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Review stage | Inputs reviewed | Key decision | Open actions | Required attendees | Approval status |
| --- | --- | --- | --- | --- | --- |
| Concept review | User needs, intended use, business case, preliminary hazards. | Proceed to planning. | Clarify home-use environment assumptions. | Sponsor, R&D, Quality, Regulatory, Clinical | Approved with actions |
| Planning review | Design plan, risk plan, standards matrix, regulatory plan. | Proceed to input definition. | Add cybersecurity threat modeling activity. | Project Manager, Quality, Regulatory, Software | Approved |
| Input review | Design input specification, traceability, user needs. | Proceed to detailed design. | Resolve battery alarm acceptance criterion. | Systems, Quality, Risk, Clinical, Service | Approved with actions |
| Output review | Drawings, software architecture, labeling, manufacturing concept. | Proceed to V&V planning. | Update IFU warnings based on risk analysis. | Engineering, Quality, Manufacturing, Labeling | Approved with actions |
| V&V readiness review | Protocols, sample size rationale, test equipment, acceptance criteria. | Start formal verification. | Calibrate test fixture before protocol execution. | V&V, Quality, Metrology, Regulatory | Approved |
| Transfer review | DMR, validation reports, training, supplier readiness. | Authorize pilot production. | Complete supplier PPAP evidence. | Operations, Quality, Supplier Quality, Regulatory | Conditional approval |

[Note] for implementation consultants: Do not close a design review simply because a meeting occurred. The record must show the evidence reviewed, the decisions made, and how unresolved items are controlled.

# 10. DESIGN VERIFICATION

Design verification shall confirm that design outputs meet approved design input requirements. Verification methods may include testing, inspection, analysis, demonstration, simulation, code review, calculation, comparative analysis, or review of validated supplier evidence. Verification protocols shall define objective acceptance criteria, sample rationale, equipment, method, environmental conditions, data recording requirements, deviations, and responsibilities before execution.

Verification failures shall be documented and dispositioned through the applicable nonconformity, deviation, issue management, design change, or CAPA process. Re-verification shall be performed when changes could affect previously verified requirements. Verification reports shall include results, deviations, statistical analysis where applicable, conclusion, unresolved anomalies, and approval.

## VERIFICATION TRACEABILITY MATRIX

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Requirement ID | Design output | Verification protocol | Acceptance criterion | Result | Report reference |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Requirement ID | Design output | Verification protocol | Acceptance criterion | Result | Report reference |
| --- | --- | --- | --- | --- | --- |
| DI-001 | DO-001, DO-002 | VTP-ECG-001 | Heart rate accuracy +/- 3 bpm. | Pass | VTR-ECG-001 Rev A |
| DI-002 | DO-004 | VTP-MECH-003 | No hazardous damage after 1.0 m drop. | Pass | VTR-MECH-003 Rev B |
| DI-003 | DO-003, DO-005 | VTP-SYS-004 | Battery alarm within 10 seconds. | Pass with deviation accepted. | VTR-SYS-004 Rev C |
| DI-004 | DO-005 | VTP-USAB-002 | Setup task support confirmed before validation. | Pass | VTR-USAB-002 Rev A |
| DI-005 | DO-003 | VTP-SW-015 | Critical alarm event logged with timestamp. | Pass | VTR-SW-015 Rev A |
| DI-006 | DO-006 | VTP-PKG-006 | Sterile barrier integrity maintained. | Pass | VTR-PKG-006 Rev B |

# 11. DESIGN VALIDATION

Design validation shall confirm that the resulting medical device conforms to defined user needs and intended use. Validation shall be performed under defined operating conditions on initial production units, production-equivalent units, or justified equivalents. Validation may include clinical evaluation, performance evaluation, simulated use, summative usability validation, packaging validation, shipping validation, shelf-life validation, software validation, installation validation, service validation, and process-output validation where applicable.

Validation protocols shall define intended use, user profiles, use environments, representative device configuration, acceptance criteria, sample rationale, data collection, adverse observation handling, deviations, and responsibilities. Validation results shall be documented, reviewed, and approved before design release or market authorization as applicable.

## VALIDATION PLAN AND EVIDENCE MATRIX

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Validation activity | User need / intended use | Population or sample | Acceptance criteria | Status | Record reference |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Validation activity | User need / intended use | Population or sample | Acceptance criteria | Status | Record reference |
| --- | --- | --- | --- | --- | --- |
| Summative usability validation | Trained nurses can set up and monitor patients safely. | 15 representative nurses. | No unacceptable use-related residual risk. | Completed | UVR-ECG-001 |
| Clinical performance evaluation | Device detects ECG signals for intended patient population. | Multisite clinical dataset. | Performance meets protocol endpoints. | Completed | CER/PER-ECG-004 |
| Packaging distribution validation | Product arrives without damage and maintains sterile barrier. | 3 lots, worst-case shipment profile. | No seal breach or product damage. | Completed | PVR-PKG-007 |
| Installation validation | Field engineers install device correctly at hospital site. | 5 simulated installations. | All critical checks completed. | Completed | IVR-SVC-002 |
| Software system validation | Final integrated system supports intended workflow. | Production-equivalent build. | No unresolved critical anomalies. | Completed | SVR-ECG-003 |
| Shelf-life validation | Accessory remains conforming for labeled shelf life. | Accelerated and real-time aging samples. | All critical attributes within limits. | In progress | SLR-ACC-010 |

[Note] for implementation consultants: Design validation is not the same as design verification. Validation must address user needs and intended use, not only specification conformance.

# 12. DESIGN TRANSFER

Design transfer shall ensure that approved design outputs are correctly translated into production specifications, inspection methods, purchasing requirements, service procedures, labeling controls, packaging controls, training materials, and release criteria. Transfer planning shall define manufacturing site readiness, equipment qualification, process validation, test method validation, supplier readiness, material availability, configuration management, DMR completeness, DHR requirements, acceptance activities, quality control plans, and post-transfer monitoring.

## DESIGN TRANSFER READINESS CHECKLIST

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Readiness item | Acceptance evidence | Owner | Target date | Status | Residual risk / action |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Readiness item | Acceptance evidence | Owner | Target date | Status | Residual risk / action |
| --- | --- | --- | --- | --- | --- |
| Device master record complete | All drawings, BOM, work instructions, test methods released. | Manufacturing Engineering | 2026-06-01 | Complete | None |
| Process validation complete | IQ/OQ/PQ reports approved for final assembly line. | Process Validation Lead | 2026-06-10 | Complete | Monitor first three lots |
| Supplier readiness confirmed | Critical supplier capability and change notification verified. | Supplier Quality | 2026-06-12 | Open | Awaiting coating supplier validation report |
| Training completed | Operators and inspectors trained on released work instructions. | Production Quality | 2026-06-15 | Complete | New hires controlled through LMS |
| Labeling and UDI released | Artwork, IFU, and UDI data approved. | Labeling Lead | 2026-06-18 | Complete | Country variants tracked |
| Post-transfer monitoring plan approved | Yield, complaints, NC and service metrics defined. | Quality Systems | 2026-06-20 | Complete | Escalate if yield below 95% |

# 13. DESIGN CHANGES

Design changes shall be identified, documented, reviewed, verified, validated as appropriate, approved, and controlled before implementation. Change impact assessment shall consider design inputs, outputs, risk management file, usability, software, cybersecurity, biocompatibility, sterilization, packaging, labeling, process validation, supplier qualification, regulatory submissions, market authorizations, clinical evidence, service documentation, inventory, distribution, and fielded devices.

The organization shall determine whether a design change requires regulatory notification, prior approval, technical documentation update, field action, customer communication, revalidation, re-verification, or post-market monitoring. Emergency changes shall be controlled through documented justification, quality approval, risk evaluation, and retrospective completion of required records.

## DESIGN CHANGE IMPACT ASSESSMENT

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Change request | Affected item | Impact assessment | Required V&V | Regulatory impact | Final disposition |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Change request | Affected item | Impact assessment | Required V&V | Regulatory impact | Final disposition |
| --- | --- | --- | --- | --- | --- |
| DCR-001 | Battery supplier change. | May affect battery life and alarm performance. | Battery life, alarm, transport test. | Assess country submissions. | Approve after V&V pass |
| DCR-002 | Software alarm algorithm update. | Affects critical alarm logic and event logging. | Software regression and system verification. | Potential notification required. | Hold release pending regulatory decision |
| DCR-003 | Label warning clarification. | Affects IFU and training material only. | Labeling review and usability rationale. | Technical file update. | Approve with controlled rollout |
| DCR-004 | Enclosure resin alternate. | Affects mechanical strength and biocompatibility. | Drop, cleaning, biocompatibility review. | Supplier and material documentation update. | Approve after risk review |
| DCR-005 | Manufacturing test fixture upgrade. | May affect final acceptance test equivalence. | Test method validation and equivalence study. | No submission impact expected. | Approve for pilot lot |
| DCR-006 | Sterile pouch seal parameter change. | Affects sterile barrier process. | Packaging process revalidation. | Assess sterilization file. | Approve after PQ completion |

[Note] for implementation consultants: Regulatory impact assessment must be performed before implementation in affected markets. Maintain evidence of rationale even when the decision is no submission required.

# 14. DESIGN AND DEVELOPMENT FILE

A design and development file shall be established and maintained for each device or device family. The file shall include or reference records that demonstrate conformity to design and development requirements, including plans, inputs, outputs, reviews, verification, validation, transfer, changes, risk management records, traceability, approvals, and unresolved issue dispositions. The file may be maintained electronically when controlled by validated document or product lifecycle systems.

## DESIGN HISTORY FILE INDEX

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| DHF section | Required content | Record system | Minimum approval | Retention trigger | Completeness check |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| DHF section | Required content | Record system | Minimum approval | Retention trigger | Completeness check |
| --- | --- | --- | --- | --- | --- |
| Planning | Design plan, regulatory plan, risk plan, project charter. | PLM/DMS | Project Sponsor and Quality | Device lifetime plus regulatory retention. | Gate 1 checklist |
| Inputs | User needs, design inputs, standards matrix. | PLM requirements module | Design Owner and Quality | Device lifetime plus regulatory retention. | Input review checklist |
| Outputs | Specifications, drawings, software, labeling, DMR references. | PLM/DMS/source control | Functional owner and Quality | Device lifetime plus regulatory retention. | Output release checklist |
| Reviews | Design review minutes, actions, approvals. | DMS | Review chair and Quality | Device lifetime plus regulatory retention. | Review action closure report |
| V&V | Protocols, reports, deviations, traceability matrix. | V&V repository/DMS | V&V Lead and Quality | Device lifetime plus regulatory retention. | V&V summary report |
| Transfer and change | Transfer report, change records, configuration baseline. | PLM/DMS | Manufacturing and Quality | Device lifetime plus regulatory retention. | DHF release audit |

# 15. TRACEABILITY AND CONFIGURATION MANAGEMENT

Traceability shall be maintained from user needs to design inputs, risk controls, design outputs, verification, validation, labeling, production controls, and design changes. Configuration baselines shall be established at defined phase gates and before formal verification, validation, regulatory submission, design transfer, and commercial release. Software builds, hardware revisions, labeling revisions, BOM revisions, manufacturing process revisions, and test method revisions shall be controlled to prevent unintended use of obsolete or unapproved configurations.

## END-TO-END TRACEABILITY MATRIX

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| User need | Design input | Risk control | Design output | Verification | Validation |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| User need | Design input | Risk control | Design output | Verification | Validation |
| --- | --- | --- | --- | --- | --- |
| Monitor patient ECG reliably. | DI-001 | RC-001 signal accuracy control. | DO-001, DO-002 | VTR-ECG-001 | CER/PER-ECG-004 |
| Warn user before power loss. | DI-003 | RC-022 battery alarm. | DO-003, DO-005 | VTR-SYS-004 | SVR-ECG-003 |
| Enable intuitive setup. | DI-004 | UR-005 usability control. | DO-005 | VTR-USAB-002 | UVR-ECG-001 |
| Maintain sterile accessory integrity. | DI-006 | PK-003 packaging control. | DO-006 | VTR-PKG-006 | PVR-PKG-007 |
| Support service investigation. | DI-005 | CY-008 audit trail control. | DO-003 | VTR-SW-015 | SVR-ECG-003 |
| Survive normal handling. | DI-002 | HZ-014 mechanical control. | DO-004 | VTR-MECH-003 | PVR-PKG-007 |

# 16. OUTSOURCED DESIGN AND SUPPLIER INTERFACES

Outsourced design activities, external laboratories, contract design organizations, software development partners, tooling suppliers, packaging suppliers, and critical component suppliers shall be controlled based on their impact on product quality and regulatory compliance. Quality agreements or purchasing controls shall define responsibilities for design records, verification evidence, validation evidence, change notification, confidentiality, intellectual property, cybersecurity, audit rights, nonconformity reporting, and record retention.

## OUTSOURCED DESIGN CONTROL REGISTER

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Supplier / partner | Outsourced activity | Required controls | Deliverables | Review owner | Acceptance decision |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Supplier / partner | Outsourced activity | Required controls | Deliverables | Review owner | Acceptance decision |
| --- | --- | --- | --- | --- | --- |
| MedTech Labs GmbH | EMC and electrical safety testing. | Approved lab, protocol approval, raw data review. | Test report and deviation log. | Verification Lead | Accepted after minor deviation closure |
| Nordic Code Systems | Embedded software module development. | Software quality plan, code review, configuration control. | Source code, unit test, anomaly log. | Software Quality Manager | Accepted with regression testing |
| SterilePack Europe | Sterile pouch design and aging studies. | Quality agreement, validation plan, material controls. | Packaging validation report. | Packaging Lead | Accepted |
| Alpine Tooling AG | Injection mold fabrication. | Tool qualification, change notification, maintenance plan. | Tooling qualification report. | Manufacturing Engineering | Accepted after dimensional study |
| Clinical Insight Ltd | Clinical evaluation support. | Protocol review, data integrity, qualification. | Clinical evaluation report. | Clinical Affairs Director | Accepted |
| SecureMed Analytics | Cybersecurity penetration testing. | Scope approval, vulnerability handling, confidentiality. | Penetration test report. | Cybersecurity Lead | Accepted after remediation |

# 17. RECORDS AND RETENTION

Records generated by this procedure shall be retained in accordance with the document and record control procedure, applicable regulatory requirements, product lifetime, and contractual obligations. Records shall be protected against loss, deterioration, unauthorized alteration, and unintended deletion. Electronic records shall be maintained in validated systems when required and shall preserve audit trails, access controls, backup, restoration, and data integrity.

Design and development plan and updates.

Design input, output, review, verification, validation, transfer, and change records.

Risk management file references and risk control traceability.

Configuration baselines, release approvals, and design history file index.

Supplier design deliverables, external test reports, and acceptance decisions.

Deviation, anomaly, nonconformity, CAPA, and regulatory impact assessment records related to design control.

# 18. EFFECTIVENESS MONITORING

The effectiveness of design control implementation shall be monitored using process indicators, quality indicators, audit results, phase gate performance, verification and validation outcomes, post-market feedback, CAPA data, complaint trends, design change quality, submission questions, transfer performance, and production yield after release. Metrics shall be reviewed by design governance and management review as applicable.

## DESIGN CONTROL EFFECTIVENESS METRICS

Use the blank table below to document organization-specific information. Maintain objective evidence, controlled record references, and approval status before releasing this template for operational use.

| Metric | Definition | Target | Review forum | Data source | Escalation rule |
| --- | --- | --- | --- | --- | --- |

Example table populated with practical entries for a large medical device manufacturer and distributor:

| Metric | Definition | Target | Review forum | Data source | Escalation rule |
| --- | --- | --- | --- | --- | --- |
| DHF right-first-time | DHF items accepted without major remediation at gate review. | >= 95% | Design quality council | DHF checklist | Escalate below target for two projects |
| Verification pass rate | Formal tests passed without unresolved major deviation. | >= 90% | Project review board | V&V reports | Escalate major failures |
| Requirements volatility | Approved input changes after verification start. | <= 5% | Program steering committee | PLM requirements log | Escalate if scope affects launch |
| Design transfer yield | First three production lots meeting release criteria. | >= 95% | Operations quality review | DHR and yield data | Escalate below threshold |
| Post-launch design issue rate | Complaints or CAPA linked to design within first 12 months. | Downward trend | Management review | Complaint/CAPA system | Escalate safety-related signals |
| Regulatory submission deficiency rate | Questions attributable to design evidence gaps. | Zero critical gaps | Regulatory review | Submission tracker | Escalate any critical deficiency |

# 19. APPENDICES AND FORMS

The following forms may be used as controlled templates or configured within an electronic product lifecycle management system. When implemented electronically, data fields, workflow approvals, audit trails, and access permissions shall preserve the intent of this procedure.

Design and development plan template.

Design input requirements register.

Design review minutes template.

Verification protocol and report template.

Validation protocol and report template.

Design transfer checklist.

Design change impact assessment form.

Design history file index and release checklist.

Traceability matrix template.

# 20. REFERENCED STANDARDS AND GUIDANCE

ISO 13485:2016 - Medical devices - Quality management systems - Requirements for regulatory purposes: https://www.iso.org/standard/59752.html

ISO 13485 overview, including risk management and risk-based decision making emphasis: https://www.iso.org/iso-13485-medical-devices.html

ISO 14971:2019 - Medical devices - Application of risk management to medical devices: https://www.iso.org/standard/72704.html

GHTF/IMDRF guidance on quality systems for the design and manufacture of medical devices: https://www.imdrf.org/sites/default/files/docs/ghtf/final/sg3/technical-docs/ghtf-sg3-n99-8-guidance-on-quality-990629.doc

GHTF/IMDRF design control guidance: https://www.imdrf.org/sites/default/files/docs/ghtf/final/sg3/technical-docs/ghtf-sg3-n99-9-design-control-990629.pdf

IMDRF/SaMD WG/N23 final, software as a medical device quality management system guidance: https://www.imdrf.org/sites/default/files/docs/imdrf/final/technical/imdrf-tech-151002-samd-qms.pdf